## Sheet1
FRAME

Captured frames are read out over usb in muliples of 1024 bytes. Each ’frame’ of data contains a HEADER, CONFIG, FRAME DATA, and enought FILLER data to fill the last USB packet. The size of the FRAME DATA is given by FRAMESIZE contained in the HEADER. Since the usb transfers are send in multiples of 1024 bytes the MAGIC fields in the HEADER can be used by drivers and applications to synchronize USB reads to the HEADER and subsequent FRAME DATA. FILLER is used to padd the FRAMEDATA to the next 1024 byte boundary. All data is transfered in BIG ENDIAN format.

	Basic Frame Layout
	Word idx	# of bytes
	0	20	HEADER
	10	108	HEADER PADDING
	64	128	CONFIG DATA
	128	768	CONFIG FILLER
	512	N	FRAME DATA
		-	FRAME FILLER

HEADER

#0 / #1	MAGIC

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	MAGICH

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	MAGICL

		MAGICH		Slways set to 0xDDDD
		MAGICL		Slways set to 0xDDDD

#2 / #3	TEMP

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	?

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	?

		?	* there should be temperature encoded here somehow. Todo- fugure out how temp is coded
		?	* for temperature monitoring...

#4 / #5	ETS DELAY

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	?

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	?

		?	ETS delay
		?

#6 / #7	ETS MIN

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	?

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	?

		?	ETS min
		?

#8 / #9	ETS MAX

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	?

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	?

		?	ETS max
		?

#10 - #127	HEADER PADDING

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	-

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

		These bytes are reserved

CONFIG DATA

The configuration data is sent in the header of every frame. It gives context for the FRAME DATA

#0	ADC_control_reg_addr

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	CRA[0:15]

		CRA[0:15]		This value should always be set to 0x000D. No idea why...

#1	ADC_control_reg_data

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					CRD[0:7]

		CRD[0:7]		Allows setting the mode. No idea what the different modes do though... you probably just want to set it to normal

		BIN	HEX	Mode
		00000000	00	NORMAL
		00000010	02	TEST

#2 / #3	VGAINA / VGAINB

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-				VGAINx[0:11]

		VGAINx[0:11]		VGAINx is used to configure the desired gain of the analog input stage for each channel. Internally a voltage between -1.8 and 1.8 volts is produced which corresponds to a VGAIN value of 0 and 4095 respectively. This voltage is fed to the variable gain amplifier wich in turn provides an analog gain between approximately 0db and 24db. The VGA itself is referenced to 0v at approximately 12db. This in combination with the ATT bit provides the total attenuation/amplification per channel. These values and graphs are approximate. Callibration is required.

		VGAINx	HEX	GAIN (DB)
		0	0000	0
		…	…	…
		1365	0555	0
		3413	0D55	24
		…	…	…
		4095	0FFF	24
		*aproximate

#4 / #5	OFFSETA / OFFSETB

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-				OFFSETx[0:15]

		OFFSETx[0:15]		Ground reference for measurments. This is a 12 bit twos complement signed value in the range of -2048 to 2047. The Ground reference is produced by a DAC and fed to the Vcom of the differential input amplifier. Positive values move the sampled values down. Negative values move samples values up.

#6	CTRL

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					ETS	ADCINT	ACA	ACB	CNGA	GNDB	ATTA	ATTB

		ETS		??
		ADCINT		Set to 1 to connect Channel 1 to both ADC inputs (and disconnect the analog input channel 2)
		ACA		Set to 1 to enable DC coupling
		ACB		Set to 1 to enable DC coupling
		CNGA		Set to 1 to ground channel A					ATTx		Rt		Rb	Vgain	db
		GNDB		Set to 1 to ground channel B					1	R10	47,00	R9	953,00	0,0470	-26,5580
		ATTA		Set to 1 to attenuate channel A					0	R16	953,00	R15	49,90	0,9502	-0,4433
		ATTB		Set to 1 to attenuate channel B					* approximate

#7	TRIGGER_MODE

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					REARM	TMOD1

		TMOD[0:1]		Trigger mode selection bits

		BIN	HEX	Mode
		00	0000	Auto	Automatically switch between Normal and continuous mode
		01	0001	Normal	A capture is made whenever the trigger event occurs
		10	0002	Single	The capture is only triggered once.
		11	0003	Continuous	Capture is triggered immediatly after the previous one. Other trigger settings are ignored

		REARM		Rearms the trigger system. Set this bit to 1 to rearm the trigger in single mode.

#8	TRIGGER_SOURCE

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					TSRC

		TSRC[0:2]		Selects the source of the trigger event

		BIN	HEX	Source
		000	0000	CH-A
		001	0001	CH-B
		010	0002	AWG-1
		011	0003	AWG-2
		100	0004	External
		xxx	xxxx	Reserved

#9	TRIGGER_SLOPE

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					TSLP

		TSLP[0:1]		Selects the slope of the trigger event

		BIN	HEX	Direction
		00	0000	Rising
		01	0001	Falling
		10	0002	Both
		xx	xxxx	Reserved

#10	TRIGGER LEVEL

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					TLVL

		TLVL[0:10]		Sets the trigger level for the trigger event. A trigger event will occur in the associated trigger mode when the sampled value passes the trigger level

#11	TRIGGERHYST

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					THYST

		THYST[0:10]		Sets the hysteresis for the trigger event. This value can be between 0 and TLVL. The lower value of the trigger is given by TLVL-THYST. The upper value is given by TLVL. Min 0. Max 1023-TLVL

#12	PRETRIGGER

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	TPRE

		TPRE[0:15]		Sets the number of samples before the trigger in steps of 1024

#13	TIMEBASE

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					TBASE

		TBASE[0:4]		Sets the timebase. This is effectively the sample rate

		BIN	HEX	Timebase
		00000	0000	2ns
		00001	0001	4ns
		00010	0002	8ns
		00011	0003	20ns
		00100	0004	40ns
		00101	0005	80ns
		00110	0006	200ns
		00111	0007	400ns
		01000	0008	800ns
		01001	0009	2us
		01010	000A	4us
		01011	000B	8us
		01100	000C	20us
		01101	000D	40us
		01110	000E	80us
		01111	000F	200us
		10000	0010	400us
		10001	0011	800us
		10010	0012	2ms
		10011	0013	4ms
		10100	0014	8ms
		10101	0015	20ms
		xxxxx	xxxx	---
		11111	001F	4ns-ETS

#14 / #15	HOLDOFF

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	HOLDOFFH

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	HOLDOFFL

	Have yet to figgure out exactly what this means... i think it is a blanking of the trigger. But i don’t know if it is before or after or … whatever

#16 / #17	FRAMESIZE

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	FRAMESIZE

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	FRAMESIZE

		FRAMESIZE		Gives the number of samples in the body of the frame. The size of the SAMPLE DATA in bytes is FRAMESIZE*4 rounded up to the next multiple of 1024

#18	AWG1-CNF0

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					AWG1EN	-				AGW1TYPE

		AWG1EN		Set this bit to 1 to enable the arbitrary waveform generator

		AGW1TYPE		Select the type of waveform

		BIN	HEX	Mode
		0000	0000	reserved
		0001	0001	custom
		0010	0002	sin
		0011	0003	cos
		0100	0004	triangle
		0101	0005	saw
		0110	0006	square
		0111	0007	delta
		1000	0008	dc*
		1001	0009	noise
		xxxx	xxxx	reserved

				*To use the generator in dc mode. The value must be set to 0 and the offset used to set the desired output voltage

#19	AWG1-CNF1

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	?	?	?	?	AWG1VOL

		AWG1VOL		arbitrary waveform generator amplitude. Unsigned

#20	AWG1-OFFSET

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-				AWG1OFFSET

		AWG1OFFSET		arbitrary waveform generator offset. Twos complement. Signed

				R = 100
				IFS = 32 * 1,25 * 4700
				Vo1 = R * IFS * AWG1OFFSET / 4096
				Vo1 = R * IFS * (4096 – AWG1OFFSET) / 4096
				Vout = 2,4 * (Vo2 -Vo1)

#21 / #22	AWG1-DELTA

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	DELTAH

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	DELTAL

	Have yet to figure out exactly how this works.

#23	AWG1-DUTY

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					AWG1DUTY

		AWG1DUTY		Duty cycle of the square wave in steps of 1/2048 %

#24 - #29
	AWG2	Identical to AWG1

#30 / #31	DIGITAL PATTERN 1

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	-

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#32	DIGITAL MASK 1

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#33 / #34	DIGITAL PATTERN 2

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	-

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#35	DIGITAL MASK 2

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#36 / #37	DIGITAL PATTERN 3

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	-

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#38	DIGITAL MASK 3

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#39 / #40	DIGITAL PATTERN 4

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	-

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#41	DIGITAL MASK 4

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#42	DELAY MAX COUNT 1

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#43	DELAY MAX COUNT 2

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#44	DELAY MAX COUNT 3

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#45	DELAY MAX COUNT 4

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				Have not documented this yet

#46	dt_StageStart; dtSerial; dtSerialCh				stage at which capture starts! (0, 1, 2, 3); parallel/serial mode; serial channel select					dt_StageStart=bit9-bit8; dtSerial=bit4; dtSerialCh=bit3-bit0

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					STAGE START					SERIAL	CHANNEL

				Have not documented this yet

#47	DIGITAL IO VOLTAGE

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					WIPER

		WIPER		This sets the digital voltage for both inputs and outputs. put formula here for approximating output voltage of the vcc_dig DAC
				VCC_DIG = …. Volts

#48	DIGITAL DIRECTION

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					DIR1	DIR0

		DIR0		Not sure but most likely 1 means output and 0 means input. DIR0 sets ios 0-5 and DIR1 sets io 6-11... probably
		DIR1

#49	DIGITAL OUTPUT WORD

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

		?

#50	DIGITAL OUTPUT MASK

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

		?

#51/#52	DIGITAL SAMPLING CLOCK PRESCALER

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	PRESCALERH

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	PRESCALERL

#53	ADC SAMPLE AVERAGING

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-					CHA	CHB	-

		CHA		Set to 1 to enable moving average of channel A					How big is the window?
		CHB		Set to 1 to enable moving average of channel B

#54 - #63

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	-

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	-

				These bytes are reserved.

#64 - #127					#192

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	CFILLERH

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	CFILLERL

		CFILLERH		Always set to 0x0000
		CFILLERL		Always set to 0xFFFF

SAMPLE DATA

The Sample data is a block of FRAMESIZE samples. Each sample contains the A channel, B channel, and digital values encoded in two 16bit words.

#0 to FRAMESIZE*2

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	SAMPLE A					SAMPLE B

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	SAMPLE B				D11	D10	D9	D8	D7	D6	D5	D4	D3	D2	D1	D0

		SAMPLE A		Channel A sample value
		SAMPLE B		Channel B sample value
		D0-D11		Digital input sample value

FILLER

USB transfers are always made in chunks of 1024 bytes. So the last chunk in the SAMPLE data is padded out with FILLER data

# 0 to next boudary of 1024 bytes

	31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16
	FFILLERH

	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
	FFILLERL

		FFILLERH		Always 0x0000
		FFILLERL		Always 0x0000
VGAINx Vs. total analog gain (dB)
VGAINx
gain (dB)